# Current Sensor

S0035 Current Sensor

Measuring range: -10~10A

Accuracy: ±2%F.S

Resolution: 0.01A

Current Sensor is used to measure the current in circuit. When current passes through resistor, it will form a tiny electric potential difference at both ends of the resistor. After circuit amplifying, it can accurately measure the current in DC circuit or low-voltage, low-frequency AC circuit.

Component

| No. | Model | Name | Picture |
| --- | --- | --- | --- |
| 1 | S0035 | Current Sensor |  |
| 2 |  | Red alligator clip |  |
| 3 |  | Black alligator clip |  |

Usage

Connect Current Sensor to any port in SenseDisc.

Connect red and black alligator clips to positive and negative poles in Current Sensor.

Series connect Current Sensor into circuit.

Calibration

Current Sensor needs to be calibrated before using.

Short connect red and black alligator clips.

Open experiment software . Click  and select ‘Current Sensor’. Click ‘Calibrate’ ‘OK’ to complete calibration.

Typical experiment

Series and parallel connection of resistor

Measure the VA characteristic curve of small bulb

Measure the electrodynamic potential and internal resistance of battery

Ohm's Law

Notes

Current Sensor should be series connected to circuit.

Current Sensor needs to be calibrated before using.

The maximum current can be applied to sensor is 10A.