# Potassium Ion Sensor

S1022 Potassium Ion Sensor

Measuring range: 0.4ppm ~ 3900ppm

Resolution: 0.3ppm

The sensor is designed with ion selective electrode. It is a kind of electrochemistry sensor which measures ion activity and concentration in solution using membrane potential. When it touches the solution with ion to be measured, membrane potential will be generated between its membrane and solution. Ion selective electrode is also called membrane electrode. There is a layer of special electrode membrane on the electrode which response to specific ion. The relationship between potential of electrode membrane and ion to be measured compliance with Nernst formula. The circuit of sensor converts potential to analog signal.

Component

| No. | Model | Name | Picture |
| --- | --- | --- | --- |
| 1 | S1022 | Potassium Ion Sensor |  |
| 2 |  | Potassium Ion Sensor electrode |  |

Calibration

Potassium Ion Sensor has been demarcated during manufacture. Normally users can use it directly.

Due to long time using, the measuring value of Potassium Ion Sensor may be wrong. If users think the obvious error occurs of concentration value, you can demarcate the sensor. The method is:

Prepare 2-5 kinds of potassium ion solution. The concentration of these solution in descending/ascending order of 10. Demarcate it successively.

Usage

- Connect Potassium Ion Sensor to any port in SenseDisc.

- Connect electrode to Potassium Ion Sensor.

Immerse the electrode into solution to collect data.

When the experiment is completed, rinse the electrode using distilled water immediately. Dry it using filter paper. Cover the electrode and preserve it in bag.

Typical experiment

Detect concentration of potassium ion in KCl.

Notes

- Immerse the electrode completely in the solution to be measured during measuring.

- When change the sample to test, rinse the electrode using distilled water first. Dry it using filter paper. Immerse the electrode into next solution. It will prevent cross-contamination.

The pH value of sample should be within 4-9. You should refer to journals as many as possible or use recognized measure method (such as national standard) to make sure proper usage.

High concentration of sodion and ammonium Ion will affect the experimental result.

Keep the plug of electrode dry and clean. Try to avoid corrosion and contamination occur on it

Try to avoid friction and strong vibration occur on electrode during using.

Do not dry the electrode using gauze or tissue. It may damage sensing membrane.