# Introduction:

The Charge Sensor can be used to measure electric quantity of a conductor. For example: it can be used to replace traditional gold leaf electroscope and electrometerin static electricity experiment.

The Charge Sensor is installed with an high precision input capacitor. The measured charge will charge capacitor so that voltage across the capacitor is increased. The device calculates voltage or charge by using potential difference between the capacitor.

It is designed with alligator clip which covered by plastic. The wires used are red and black (for ground).

There is a button equipped on the sensor which is used to change the measuring range. Press the button and the measuring range will become high otherwise it will be low.

There is a reset button equipped on the sensor which is used to discharge the capacitor.

Usage:

Connect alligator clips to Charge Sensor. Connect Charge Sensor to datalogger. Select proper measuring range. The internal resistor designed in Charge Sensor is extremely high. It will gain charge from the surrounding environment. Press the reset button before experiment. Press it again to discharge the sensor after collecting.

In order to eliminate the effect from static electric field, ground the sensor during using. If it can not be grounded, user should stand on ground and keep still. For those experiments which need high precision charge measuring, fix the sensor to ensure the result is accurate. Be careful of effect from weather during using. The air humidity can also affect the accuracy of data measuring. Dry the surrounding air to ensure the accuracy.

Calibration:

The reset button of TS2119 Charge Sensor must be pressed before every experiment.

Typical experiment:

- Compare charge quantity and symbol of different objects

- Static electric phenomenon

- Charge sharing between conductors

- Faraday-ice bucket experiment

- Electrostatic screening

- Measuring charge quantity and symbol

- Charge density on surface of conductor

# Experimental case:

Measuring charge quantity and symbol:

Product General Assembly

Experimental procedures：

- Connect experiment apparatus. Calibrate Charge Sensor.

- Open SWR iLab v8.0. Click ‘new’ to create new experimental template.

- Click ‘new page’ to add new folder.

4. Click ‘add line’ to add ‘Charge-Time’ coordinate. Set proper experiment time and acquisition interval.

5. Connect black alligator clip to ground.

6. Click ‘start’ to collect data. Connect rubber bar which has been rubbed with fur to red alligator clip.

7. Observe and analyze the experimental result.

8. Clear up experimental apparatus.

Charge Change with Respect to Time

Notes:

The wire used to connect alligator clips to sensor must be covered by plastic.

Discharge the built-in capacitor before any experiment. Short connect red and black alligator clips or press reset button.

For those experiment which need high precision charge measuring, fix the sensor to ensure the result is accurate.

The air humidity can also affect the accuracy of data measuring. Dry the surrounding air to ensure the accuracy.

# Text boxes

Charge Sensor

Product No.: TS2119

Low range: -22nC～+22nC

Resolution: 0.01nC

Accuracy: ±2%

High range: -220nC～+220nC

Resolution: 0.1nC

Accuracy: ±2%

Address: No.14 Junnong Road, Qinhuai District, Nanjing, China.

Website: //www.sinoswr.com

Post Code: 210007

Tel: +86-400-828-8387