Introduction:

PH Sensor electrode is made of internal reference electrode and glass electrode. It is used to measure the hydrogen ion concentration of a solution and display the pH value of the solution.

Usage:

Unscrew the pH electrode from protective bottle (with saturated KCl solution) before using. Check if there is air bubble in the glass ball. Swing to eliminate it. Try not to touch other objects.

Connect the electrode to PH Sensor. Connect PH Sensor to datalogger. The electrode should be cleaned up by using deionized water or distilled water. Its surface should be dried by using tissue. Demarcate and immerse it into liquid to be measured. Wait a while until the reading become stable and analyze the data.

Demarcation:

Before measuring, it is required to demarcate the pH electrode. When demarcating, the standard buffer solution with the pH value similar to that of the target solution should be used. The pH buffer solution can be prepared by buying buffer solution reagent from the chemical reagent shops or by referring to the standard preparation method. For the details to demarcate the PH Sensor via software, refer to software operating instruction:

1. Clean the electrode by using deionized water or distilled water. Rinse the electrode using standard pH 4.00, pH 6.86 or pH 7.00 solution. Immerse it into the standard pH 4.00, pH 6.86 or pH 7.00 solution. Open the software and select the demarcation interface. Select channel to display the pH value. Click the ‘Demarcate’ button. Enter corresponding pH value at the L-Value box and click ‘Demarcate’ button. See the Figure as follows, input ‘4.00’.

Take the electrode out and rinse it using deionized water or distilled water. Use the other kind of standard solution to rinse it. Immerse the electrode into the solution. Enter corresponding pH value at H-Valve box, as shown in the Figure above. Input is ‘9.18’. Click the ‘Sure’ button.

Take the electrode out and rinse it using deionized water or distilled water. Use the testing solution to rinse it and it will be ready to measure the testing solution.

Typical experiment:

Acid-base neutralization drop

pH value of different brine solution

Acidity of phenol

Mechanism of organism maintaining pH stability

Explore impact of pH value on pectinase activity

Experimental case:

Acid-base neutralization drop

Product General Assembly Diagram

Experimental procedures:

- Connect apparatus as shown above.

- Pour diluted hydrochloric acid into drop tube and NaOH into beaker.

- Open SWR iLab and create ‘Acid-base neutralization’ template.

- Open syringe control valve to let the solution drop at an appropriate speed. close control valve and stop the experiment.

- Calculate the derivation for the curve and select the cross-reference line. Find titration endpoint from catastrophe point. Read the volume V1 of the acid solution consumed when the liquid gets completely neutralized. Calculate the concentration of sodium hydroxide (c2) based using formula c1*v1=c2*v2.

Acid-Base Neutralization Drop

Notes:

1. Rinse the electrode before every measuring. Do not wipe its glass ball using tissue. Otherwise it may cause damage to the electrode. It is better to clean the electrode using the testing solution.

2. Keep the glass ball of the pH electrode wet in order to maintain the ion exchange process. It is required to preserve the glass ball in the KCI solution. Check the KCI solution’s volume periodically. If the electrode dries, immerse it into the KCI solution for 2 hours to make it wet.

3. Protect the glass ball of the pH electrode. Do not knock, collide or rub it. Otherwise it may cause damage to the electrode.

4. Keep electrode clean and dry when connecting and unplugging it. If not, it may lead to its insulation failure and make the result not accurate or make the experiment unsuccessful.

5. We have found that the conductivity sensor has mild interference to this product. Using other devices with it will not cause this problem.

6. The pH electrode has limited service life, which mainly depends on its protection and testing solution. Do not store and preserve the electrode in the distilled water or the deionized water. Otherwise it may cause the saturated solution on the electrode to move. Never use the pH electrode in the extreme acid-based or alkaline-based solution or at the extreme temperature.

7. Do not use the pH electrode in the absolute ethyl alcohol, concentrated sulphuric acid and other dehydration media since they may damage the hydration gel layer on the glass membrane surface. The pH electrode can not either be immersed in the neutral or the alkaline buffer solution. Long-term immersion in such solution may cause the pH glass membrane insensitive to respond.

8. The material of PH electrode is PC. It will dissolve in some kinds of solution such as ethyl acetate, carbon tetrachloride, Cross-linked polythene and tetrahydrofuran. Those solution will damage the cover of electrode. You should use the electrode covered by glass.

9. After long term using, the rate of data collecting will drop. Immerse the measuring side of electrode in 4%HF for 3-5s or diluted HCl solution for 1-2 minutes. Rinse it using distilled water and immerse it in potassium chloride (4M).

10. Common contamination and detergent:

| Contamination | Inorganic metal oxide | Organic fats and oils | Resin polymer | Protein blood deposit | Pigment material |
| --- | --- | --- | --- | --- | --- |
| Detergent | <1 mol/L diluted acid | Diluted detergent (alkalescence) | Dilute alcohol acetone, diethyl ether | Acidity enzyme solution | Diluted bleach solution, hydrogen peroxide |

11. This product can automatically compensate temperature. User does not need to make temperature compensation.

# Text boxes

Address: No.14 Junnong Road, Qinhuai District, Nanjing, China.

Website: //www.sinoswr.com

Post Code: 210007

Tel: +86-400-828-8387

PH Sensor

Product No.: TS2202

Measuring range: 0～14

Resolution: 0.003

Accuracy: ±0.1