Introduction:

Spirometer is used to measure the amount of the air that is inbreathed, breathed out and the rate that air is expelled from the lung. In anyone’s breath cycle, the amount of air inbreathed and breathed is called tidal volume. It is about 0.5L. There is extra amount of air breathed after normal breathing, which is called intake reserve and exhalation retention.

There is a semiconductor micro pressure sensing component in Spirometer. It is able to sense pressure change within ±35kPa. There is an air cavity on rubber hose to detect airflow rate of human breathing. Sensor senses small pressure difference formed when air flows through it. The pressure difference is direct proportional to the rate of air flow. By using the software, it can transfer pressure value to air flow rate. Eventually, it will be transfer to air flow volume.

Usage:

1. Assemble the Spirometer and connect it to datalogger and computer. Open SWR iLab to create new experiment, new page and set experiment time and acquisition interval.

Spirometer

2. Sit regularly or stand still when using the product. Do not look at the data being collected if possible. It will not change the way you breath subconsciously.

3. Attach the airflow head to your lips.

4. Pinch your nose to make sure all the air breathed out coming from your mouth if necessary (normally not necessary except the result is qualitative).

5. Breath calmly in order to get used to the sensor when breath becomes stable. The gas exhaled from lung is positive airflow and the gas inhaled is the negative airflow.

6. During using interval, place the Spirometer on the surface of workbench. Try to avoid the heat of your hand affecting the result.

Typical experiment:

Measure the respiratory rate of human

Measure vital capacity

Experimental Case:

Measure vital capacity

Product General Assembly

1. Connect apparatus as shown above. Assemble the filter, airflow header and this sensor before using. Connect sensor to datalogger and datalogger to computer.

2. Open SWR iLab to create new experiment. Add ‘Respiratory Rate-Time’ coordinate. Set experiment time 5min and acquisition interval 100ms.

3. Click ‘Start’ to collect data. Attach the airflow head to your lips and breath out at one time. Let all air flow coming from your mouth only.

4. After breathing, click ‘Stop’. The collected curve is respiratory rate changing with respect to time.

5. Click ‘Selection’ to select the area where you have breathed. Click ‘Curve manager’ and choose ‘Time-Respiratory Rate0’. Click ‘Integration’ in pull down menu of ‘Analysis’. The shaded area in figure is the integral area. Value shown in the shaded area shows the vital capacity.

Experimental Result (Integrated Value is Vital Capacity)

Notes:

1. The product can only be used for teaching. Follow the local safety behavior policy when using this product. The instructor has responsibility to confirm the user has the medical history of respiratory or cardiovascular and cerebrovascular diseases.

2. The airflow header provided contains a filter, which can filter most of virus, fungus and bacteria.

3. Every student should only use their own airflow header to prevent cross-infection.

4. The fixed airflow header is not required to replace every time as long as the outlet is protected by the airflow header. If the fixed airflow header is used carelessly by someone, it should be replaced immediately.

5. When using the Spirometer, the tester should not be nervous. Teacher should make sure students under testing are healthy and has no medical history of the cardiovascular disease, or has no respiratory problems (such as asthma). The Spirometer may possibly make some testers feel uncomfortable when it is used. Calm testers as much as possible. If the tester feels really uncomfortable, stop experiment immediately.

6. Make sure the airflow head is well sealed when it touches to user’s lips. Keep your tongue flat so that the product will not mess up with ventilation opening.

7. Cleanliness: The outer casing can be cleaned by a soft and wet cloth and the cloth should be painted with the disinfectant at the standard medical level or the Milton hypochlorite solution (refer to the explanatory packaging on the bottle). Do not use alcohol. Otherwise it may break the plastic.

8. In any cases, do not immerse the Spirometer or other airflow header into any cleaning solution. Do not disinfect or sterilize it with hot water, steam or hot air.

# Text boxes

Address: No.14 Junnong Road, Qinhuai District, Nanjing, China.

Website: //www.sinoswr.com

Post Code: 210007

Tel: +86-400-828-8387

Spirometer

Product No.: TS2211

Measuring range: -10L/s～10L/s

Resolution: 0.01 L/s