Introduction:

The sensor is designed with ion selective electrode. It is a kind of electrochemistry sensor which measures ion activity and concentration in solution using membrane potential. When it touches the solution with ion to be measured, membrane potential will be generated between its membrane and solution. Ion selective electrode is also called membrane electrode. There is a layer of special electrode membrane on the electrode which response to specific ion. The relationship between potential of electrode membrane and ion to be measured compliance with Nernst formula. The circuit of sensor converts potential to analog signal.

Usage:

1. Take out the cover of electrode. Rinse the electrode using distilled water. Dry it using filter paper.

2. Connect electrode to sensor, datalogger and computer.

3. Open SWR iLab. Calcium Ion Sensor has been calibrated before manufacturing. Immerse the electrode into solution. Collect data when the electrode becomes stable.

4. When the experiment is completed, rinse the electrode using distilled water immediately. Dry it using filter paper. Cover the electrode and preserve it in bag.

Typical experiment:

Detect the concentration of calcium ion in milk

Notes:

1. To maximize the service life and best function, it is suggested that preserve the electrode in fridge and temperature around 5℃.

2. Immerse the electrode completely in the solution to be measured during measuring.

3. When change the sample to test, rinse the electrode using distilled water first. Dry it using filter paper. Immerse the electrode into next solution. It will prevent cross-contamination.

4. The pH value of sample should be within 4-9.

5. Keep the plug of electrode dry and clean. Try to avoid corrosion and contamination occur on it.

6. Try to avoid friction and strong vibration occur on electrode during using.

7. Do not dry the electrode using gauze or tissue. It may damage sensing membrane.

8. Standard calcium ion solution should be 2-5 kinds. The concentration should be progressive increase or decrease by 10 times.

9. The range of standard solution should contain the concentration of solution to be measured.

# Text boxes

Address: No.14 Junnong Road, Qinhuai District, Nanjing, China.

Website: //www.sinoswr.com

Post Code: 210007

Tel: +86-400-828-8387

Calcium Ion Sensor

Product No.: TS2232

Measuring range: 0.4-4000ppm

Resolution: 0.7ppm

Accuracy: ±1.5%