Introduction:

CO2 Sensor is infrared gas absorption and testing sensor. It has good selectivity, high sensitivity, oxygen free dependent, long service life, low power consumption and no preheating requirement.

Calibration:

Due to long-term continuously use, the CO2 Sensor’s measuring value may have deviation. You can calibrate it when obvious error occurs.

The calibration method is:

Calibrate it when the sensor is placed in clean air (Sunshine is good with free flow air outdoor. The CO2 value is commonly between 350 to 450 ppm). When output concentration is stable, its value has great difference with actual concentration (such as: customer thinks the actual concentration is about 400 ppm, but the measuring value is 1000 ppm). Manual calibration function is to calibrate the measuring value to the program default value 400 ppm. The calibration method is: the output concentration of the testing sensor in the clean air is stable (wait for about 5 minutes). Use pinpoint hold against the sensor reset button for 5 seconds. Generally, reset is not required.

Usage:

CO2 Sensor can be used to detect concentration of CO2. There are ventilation opening on the electrode wall. Connect sensor to datalogger and computer. Place the electrode into the air to be measured.

Typical experiment:

Plant's photosynthesis

Explore breathing regulation of saccharomycetes

Assessment of air quality

Generation velocity of CO2 gas in chemical reaction

Compare gas exhaled by human with the environmental gas

Experimental Case:

Compare gas exhaled by human with the environmental gas

。

Product General Assembly

Experimental procedures:

1. Connect apparatus as shown above. Assemble the electrode. Connect O2 sensor and CO2 Sensor to datalogger. Connect datalogger to computer.

2. Open SWR iLab to create new experiment. Click ‘add line’ to add ‘O2-CO2’ coordinate. Set experiment time 1min and acquisition interval 100ms.

3. Click ‘Start’. Place oxygen and CO2 Sensor’s electrode at the outlet of airflow head respectively. Exhale air and then observe the change of the content of oxygen and CO2 acquired by the software.

Compare Gas Exhaled by Human with the Environmental Gas

Notes:

1. Calibrate the sensor periodically which should not be greater than 3 month.

2. Impact due to temperature and humidity to the sensor is very small. It can be used properly under condition where relative humidity is 95%.

3. Try not to impact testing result by exhaled CO2.

4. If liquid flows into the electrode, it will destroy the sensor completely. Try not to use the sensor under condition which is full of dust.

5. CO2 is heavier than air so that it will sink to bottom of bottle. Use container as small as possible when doing experiment because it will be easier to collect CO2.

6. Reserve the electrode in a place without CO2 due to the working principle of electrode. When it is placed inside room, it can be stored by 2 years.

7. The CO2 Sensor is an infrared gas absorption detector named infrared CO2 Sensor. It has good selectivity, high sensitivity, no oxygen dependence, long life, low power consumption without preheating by comparing with the chemical sensor.

Additional remarks for the sensor resolution:

| Measuring range | Resolution |
| --- | --- |
| 0～2000ppm | 1ppm |
| 0～5000ppm | 1ppm |
| 0～8000ppm | 1ppm |
| 0～1%VOL | 1ppm |
| 0～5%VOL | 10ppm |
| 0～10%VOL | 10ppm |

Additional remarks for the sensor accuracy:

| Measuring range | Accuracy |
| --- | --- |
| 0～1000ppm | ±60ppm |
| 0～1500ppm | ±95ppm |
| 0～2000ppm | ±125ppm |
| 0～3000ppm | ±175ppm |
| 0～5000ppm | ±280ppm |
| 0～50000ppm | ±3000ppm |
| 0～100000ppm | ±8000ppm |

# Text boxes

Address: No.14 Junnong Road, Qinhuai District, Nanjing, China.

Website: //www.sinoswr.com

Post Code: 210007

Tel: +86-400-828-8387

CO2 Sensor

Product No.: TS2206

Measuring range: 0～100000ppm

Resolution: ≤10ppm

Accuracy: ≤±8000ppm

Reset Button