Introduction:

Nitrite Ion (NO2-) is a kind of toxic atom radical. Nitrite consists of NO2- and metal. It is a kind of white opacity crystal and looks like salt. Industrial salt is made of chemical materials which contains lots of nitrite. For adults who take 0.2~0.5g of nitrite may cause death due to poison. It is also a kind of cancerogenic substance. According to scientific study, esophagus cancer can be caused by amount of nitrite intake.

Nitrite Ion Sensor is a kind of electrode sensor which measures concentration in solution. It detects the concentration in real time using datalogger and experiment software. The electrode is designed with solid PVC polymer membrane. It is used to measure concentration of nitrite ion in solution and suitable for outdoor using.

Usage:

Nitrite Ion Sensor needs to be demarcated before using. You need to wait until the data reading become stable. The demarcate method is shown below:

Connect Nitrite Ion Sensor to SenseDisc. Connect SenseDisc to computer. Clean and dry the probe and immerse it into standard nitrite solution ‘1’. Wait 5 minutes until the reading becomes stable. Insert first value in software (the concentration of nitrite solution ‘1’ in mg/L). Click demarcate. Use standard nitrite solution ‘2’ to demarcate second value (same as ‘1’). Demarcate is completed.

Configuration of standard solution:

1. 1M nitrite- standard solution (for attenuation)

Use analytical balance to measure and take 6.9g sodium nitrite which has been dried for 2 hours under 105℃ (analytically pure). Dissolve it in pure water and attenuate to 100mL. The standard 1M Nitrite- solution is made.

2. 0.1M nitrite- standard solution (460mg/L)

Use transfer pipette to take 10mL 1M nitrite- standard solution and pour it into 100mL volumetric flask. Attenuate it to 100mL using pure water. The 0.1M nitrite- standard solution is made.

3. 0.01M nitrite- standard solution (46mg/L)

Use transfer pipette to take 1mL 1M nitrite- standard solution and pour it into 100mL volumetric flask. Attenuate it to 100mL using pure water. The 0.01M nitrite- standard solution is made.

4. 0.001M nitrite- standard solution (4.6mg/L)

Use transfer pipette to take 10mL 0.01M nitrite- standard solution and pour it into 100mL volumetric flask. Attenuate it to 100mL using pure water. The 0.001M nitrite- standard solution is made.

Configuration of internal pipette solution:

The sensor requires internal pipette solution. When it is finished, you need to inject again. Inject it from the side hole on probe. It is 3mol/L KCl solution. The configuration is dissolving 223.5g KCl in beaker and pour it into 1L volumetric flask. Fill it to 1L.

Notes:

To maximize the service life and best function, it is suggested that preserve the probe in fridge and temperature around 5℃.

When measuring concentration of nitrite in solution, you should configure two kinds of nitrite solution to demarcate. Try to configure nitrite solution which has same concentration as the solution to be measured in order to increase accuracy.

Stable time: at least 5 minutes. Best PH range: 4.5-8. Temperature range: 5℃—50℃. Best temperature: 25℃.

The following ion will interrupt nitrite (the coefficient is shown in bracket): acetate (0.001), fluorinion (0.0008), chloridion (0.00005), nitrate (0.00001), sulfate radical (0.00001). The error will occur only when concentration of these matter is several times of nitrite. For example, when the concentration of chloridion is same as nitrite, the effect is only 0.005%. To make the effect up to 5%, the concentration should be 1000 times of nitrite.

The effect of cyanide is very significant. It is forbidden to exist or only accepted to exist when its concentration is far lower than nitrite.

The electrode filling liquid should be injected before using every time. It is suggested to discard it if the sensor will not be used in short period of time. Otherwise it will evaporate.

# Text boxes

Address: No.14 Junnong Road, Qinhuai District, Nanjing, China.

Website: //www.sinoswr.com

Post Code: 210007

Tel: +86-400-828-8387

Nitrite Ion Sensor

Product No.: TS2216

Measuring range: 4.6ppm～4600ppm

Resolution: 0.1mg/L

Accuracy: ±1.5%FS